Replacement HP 8601A Output Amplifier Module Service Manual

© Matthew D’Asaro 2012

www.6540rom.com

Disassembly for Service:

Unlike the original module, which is a sealed unit, this replacement can be repaired if trouble ever arises. However, proper procedure and care must be taken when disassembling the unit to avoid damage. A vacuum desoldering system is highly recommended.

- Remove the output amplifier module from the 8601A by pulling it straight up out of its socket. Note that the pins are easily broken if bent.

- Using a vacuum desoldering system (Hakko 472 or equivalent recommended) remove the solder from each of the points indicated in Figure 1. Take extra care to thoroughly heat pin 8 (indicated with a blue arrow) as it is connected to the ground plane and thus will require extra heat.

- Carefully pry the lid from the mainboard with a small screwdriver. If they do not separate easily DO NOT FORCE THEM APART. Rather, repeat the desoldering procedure (adding fresh solder before using the desoldering system again often helps) and then try to pry them apart again.

Once the lid is removed, the unit is easily serviced. A schematic, board layout diagram, and parts list are provided on the following pages. Note also that the module will perform correctly with the lid removed, so the lid should only be reinstalled after it has been verified that the module is operating correctly.

Reassembly:

- Using the desoldering station, remove all excess solder from the tops of the pins and the bottom of the lid. When put in place, the lid should rest firmly on the tops of the transistors.

- Apply a small amount of thermal grease to the top of each transistor.

- Use a small c-clamp to firmly clamp the lid onto the module.

- Solder each of the points indicated in Figure 1.

- Reinstall the module into the 8601A by pressing it firmly into place.

Figure 1: Solder joints holding the lid of the module in place.